V Final Term Paper: KELOS Endovascular Stent

Chih-Han (Matthew) Chen, Ryan Potechin, Mark Rotering, Faith Toner, & Megan Wirtjes

- Definition of Unmet Clinical Need

  - Abdominal Aortic Aneurysm (AAA)

An abdominal aortic aneurysm occurs when the aortic walls weaken and cannot control the flow of blood at the same rate throughout the aorta. This change in rate results in blood pooling in that section of the aorta and the walls of the aorta expanding. To be considered an aneurysm the aorta section must expand to a diameter one and one-half times greater than the original diameter.[1] At the renal arteries, this is an increase from a 2.0 cm diameter to a diameter of 3.0 cm. Risk factors for aortic aneurysms include age, sex, and a history of smoking. If the aneurysm continues to grow without intervention the aortic walls could become extremely weak or the pressure with the walls could mount to the point of an aortic rupture. Ruptured AAAs are attributed to causing 4%-5% of sudden deaths in the United States. Aneurysm size and expansion rate contribute highly to the possibility of rupture.[2]

  - Endovascular Aneurysm Repair (EVAR)

In order to reduce the chance of an AAA rupturing an endovascular stent is often used to facilitate the blood flow through the aorta. This stent provides a manufactured wall for the aorta to reduce the effects that weakened aortic walls have on the pooling of blood that originally occurred. These stents do not decrease the size of the AAA once they are placed, they simply prevent the AAA from getting bigger and possibly rupturing. An endovascular stent is placed via a catheter wire through the femoral artery. Using X-ray and ultrasound imaging techniques the surgeon guides the wire into the aortic aneurysm. Once the catheter is in place the stent is guided over the catheter and expanded into place. This process creates a seal at both the proximal and distal ends of the aneurysm and facilitates blood flow through the stent and along the path of the aorta. Complications from this procedure include endoleaks, infection, and stent migration.[3]

  - Endoleaks

Endoleaks occur when there is some defect in the stent, stent placement, or collateral vessels that causes blood flow to not be directed into the stent as intended. There are 5 different types of endoleaks; Types 1, 2, 3, 4, and 5. Type 1 endoleaks occur when the seal at the attachment type is dislodged causing leakage. Type 1 endoleaks are further broken down to indicate which site attachment has been compromised. Type 1a refers to the proximal site attachment, Type 1b refers to the distal site attachment, and Type 1c refers to a contralateral non-graft iliac artery. Type 2 endoleaks occur when there is leakage through a collateral vessel. Type 3 and Type 4 endoleaks both are due to issues with the stent, either mechanical failure or graft porosity issues respectively. Type 5 endoleaks are attributed to aneurysm expansion when there is no leakage.[4]

    - Type 1a Endoleaks

Type 1a endoleaks occur when the seal between the endovascular stent and the proximal site attachment is compromised. This allows an influx of blood to flow directly into the aneurysm sac and can cause the aneurysm to grow at an exponential rate. Type 1a endoleaks are exhibited in approximately 10% of all EVAR cases and are known to have the highest rate of rupture of any other endoleak type. Due to the possibility of a rapid influx of blood into the aneurysm sac Type 1 endoleaks have been regarded as difficult to manage and prevent rupture.[5]

- Design Solutions

  - Design

There are many components that go into making a functional stent. There is the general stent structure, the metal pattern, the wall pattern, and more. The design for the KELOS Endovascular Stent combines a traditional approach and a novel design.

The basic design is quite traditional in that it uses materials and designs common to endovascular stents on the market. The structure of the stent is made of nitinol formed in a zig-zag pattern. This material and pattern provide high strength and flexibility for the stent allowing the stent to conform to the blood vessel while also providing support [6]. The walls of the stent are also traditional since they are made of ePTFE. This material is widely used and is flexible, durable, smooth, and biocompatible making it ideal for stents [7].

The novelty of the KELOS stent comes in the wall pattern used. The KELOS stent takes a new approach by implementing a gill pattern into the ePTFE that makes up the walls. This design allows for reintroduction of blood from the aortic sac back into the main vessel without impairing normal flow. Overall, this design helps create a safeguard against rupture of the aortic sac by providing an outlet when pressure builds up enough to open the gills.

Figure 1: Schematic drawings of the KELOS Endovascular Stent

  - Materials

The materials considered in the design process stem from industry standards of biocompatible endovascular stents. The composition of endovascular stents include material used for the skeletal frame, usually metal or metal alloys, and material used for the graft that focuses on flexibility and serves as a tube-like structure to guide blood flow in abdominal aneurysms. For both components, several characteristics must be fulfilled to achieve the desired outcome of endovascular aneurysm repair (EVAR) operations.

    - Skeletal Material

The metal frame of an endovascular stent provides rigid structure yet enough flexibility for ease of insertion through blood vessels. It must go through compression and return to the original structure upon deployment at the target site. The choice of metal materials comes from their inherent ductility, or the ability of a material to be stretched, pulled, or drawn into a thin wire or thread without breaking. Its high strength provides a structure rigid enough to withstand vascular pressure from the walls to retain shape. Important considerations include the elastic modulus, stiffness, as well as its Poisson’s ratio. Furthermore, biocompatibility is certainly essential for long-term implantation to mitigate any potential toxicity, corrosivity, and allergies. The industry standard for endovascular stents are stainless steel and nitinol. [8]

The advantages of stainless steel as the choice of material are due to its high stiffness and strength at 192 GPa. It allows for good retainment of structure with amazing mechanical properties. Stainless steel is corrosion resistant and durable for a very long time. [[9]] Specifically used in abdominal vessels, stainless steel has been studied and shown to be advantageous in reducing initial thrombosis. However, it has also been shown to cause complications in late-stage chronic inflammation and neointimal hyperplasia upon long-term implantation. Disadvantages of stainless steel include corrosiveness and high density. Again, corrosiveness contributes its potential downfall in long-term implants especially for younger patients.

Nitinol is a metal alloy of around equal nickel and titanium developed at the Naval Ordnance Laboratory in 1965. It is an alloy with atypically high elasticity for up to 10% strain without yielding. Normal metallic materials can only endure around 0.2% to 2% strain before yielding and permanently deforming. [10] Furthermore, Nitinol has temperature dependent shape memory, where undergoing heating allows Nitinol to return to its original shape after deforming. [11] The downside of Nitinol includes high cost, and potential Nickel toxicity in human bodies. The small size of endovascular stents usually mitigates devastating effects of Nickel toxicity but should still be considered. For its flexibility, Nitinol is implemented as the structural metal for the KELOS design.

    - Graft Material

The graft material of endovascular stents contributes to retaining the tube shape that guides blood flow through the aneurysm sack. Desired properties include substantial rigidity and strength to hold blood flow and prevent expansion to harm the arterial walls but are flexible enough for insertion and varying vessel geometries. Similar to the structural material, the graft also needs to have a low friction coefficient and high chemical and temperature resistance. Industry standards for graft materials include polyethylene terephthalate (PET/Dacron) or expanded polytetrafluorethylene (ePTFE). [12]

PET is strong, lightweight, and has great mechanical properties. It provides a more rigid structure compared to ePTFE due to its innate chemical composition. However, due to its hardness, it needs to be woven into a fabric-like material to create flexible grafts. [13] Furthermore, it does not have great biocompatibility so pre-clotting or coating is needed to contain host blood pre-insertion.

The second option is ePTFE, with advantages of having low friction coefficient, good chemical resistance, and high temperature resistance. [13] It is very compliant which suits the insertion process and ensures safety of vessel walls from being damaged. Compared to PET, this material is less rigid, which can be an issue in designing geometrical properties on the novel invention. For its wide use and proven efficacy for stent insertion, ePTFE is chosen for our target graft material for endovascular stent design.

- Background

  - Due Diligence

    - Tissue Responses to Endovascular Stents

The major common tissue responses caused by endovascular stents include thrombosis, restenosis, and neointimal hyperplasia. Thrombosis causes vessel closure and obstruction due to blood clots composed of platelets and red blood cells. Thrombosis may be caused by various problems, such as under expanded stent or small vessels, causing stasis and turbulence flow at the insertion site. [14] Delayed healing of insertion may also cause endothelial injury and disruption of blood flow. Restenosis refers to the re-narrowing of the artery at the aneurysm site, causing fatty tissue to reduce the vessel diameter. Restenosis may happen to any patient, forming scar tissue under stent lining. [15] Lastly, neointimal hyperplasia develops neointimal tissue caused by internal body response, producing smooth muscle cells and scar tissues at the site of insertion, also narrowing the vessel. [16] Modern treatments of these issues include drug-eluding stents, antiplatelet therapy, or unfortunately repeated angioplasty or coronary intervention.

    - Structural Issues with Endovascular Stents

Structural issues that require consideration include stent migration, stent fracture, and kinking. Stent migration is the result of pulsatile force and oversizing of grafts. It can move to unwanted places and cause a series of issues depending on where it ends up. It is responsible for most late-stage complications and occurs up to 28% of stent implants. [17] A stent fracture is caused by overloading and fatigue, usually due to exceeding the tensile stress limit of the chosen material. It is less likely to occur, at about 1% to 8% of endovascular stent implants. However, fracture may be the most lethal of failure cases. Stent kinking happens in approximately 4% of stent patients and it causes the most endoleaks. The cause of kinking can usually be identified by graft migration and imperfect insertion of stent. [18]

  - Design Process

To design the KELOS stent the team started by simplifying the problem of type 1a endoleaks to its most basic form. This problem was defined as the need to keep blood from filling up in the aortic sac to the extent that it ruptures. From there the team decided to focus on designing a stent that can allow blood to flow from the aortic sac into the main vessel without flow in the opposite direction.

To solve this problem the KELOS team employed the concept of biomimetics. In searching for similar concepts in nature the team was able to identify similar concepts including the gills on a fish. During the research phase, the team found that gills on a fish act as one-way valves allowing water to pass into the fish through the mouth and out of the fish through the gills. It was also found that gills help prevent the flow of water back into the fish through the gills [19], [20].

Understanding this the KELOS team decided to implement this concept into the design of the KELOS stent. By incorporating a gill structure into the walls of the stent the KELOS team was able to achieve the goal of allowing blood to flow into the stent but not out of it. Overall, this allowed the team to design a stent to help solve the problem of type 1a endoleaks by preventing rupture of the aortic sac in case of leakages.

  - Human Factors

Insertion of the KELOS Endovascular Stent should be the same as traditional stents that are currently on the market. A small incision will be made in the iliac artery where the catheter and stent will be threaded to the site of the aneurysm. There, a vascular surgeon will ensure proper alignment, and will fix the graft using the pronged fixation device that is isolated from the sealing mechanism. After the stent is fixed, the surgeon will manually inflate a balloon to expand the stent to full size. The surgeon can then adjust the sealing mechanism to best fit the patient’s anatomy. Having an adjustable sealing mechanism that is independent from fixation ensures a tight seal with no early endovascular leaks while also allowing for use in the most diverse anatomies due to varying aneurysm neck lengths and the presence of tributary vessels.

Since the goal of the device is to mitigate the risk of rupture due to secondary endovascular leak type 1, the difference between the human factors of our device occurs within follow up procedures. Traditionally, a type 1 endoleak leads to the highest risk of aneurysm rupture, and often, death.  The KELOS Endovascular Stent allows for unidirectional blood flow from the aneurysm sac to the lumen of the stent which extends the life of the defective stent until a vascular surgeon can fix the defective seal or replace the stent entirely. Stents that are currently on the market that do not adapt to the changing anatomies and physiologies of the aneurysm sac do not allow for as much or any time to replace the damaged stent due to type 2 endoleak.

  - Competition

Currently, there are no endovascular stents on the market that look to mitigate endovascular leak type 1 by reintroducing aneurysm sac blood to the stent lumen. However, some devices that were used to draw inspiration from are the Alto Abdominal Stent Graft System by Endologix and the Gore Excluder AAA Endoprosthesis. The Gore Endoprosthesis features a modular sealing mechanism at the proximal end of the stent and has achieved no instances of early type 1 or 3 endoleak, migration, rupture, or stent fracture after a one-year follow-up study [21]. While the Excluder performs well and shows no early endoleak, the delayed risk of secondary endoleak due to a sealing issue caused by changes in aneurysm anatomy or physiology is not addressed and rupture is still a concern. The Alto system also has an adaptive sealing mechanism featuring an integrated balloon system which, like the Excluder, performs well at preventing early endoleak upon insertion [22]. It also includes a separation of fixation and sealing mechanisms which increases the patient population the stent can be used in. The Alto system also fails to adapt to changing pressures within the aneurysm sac and prevent rupture via endoleak type 1. The KELOS Endovascular Stent would be the first of its kind on the market.

  - Regulatory

According to the U.S Food and Drug Administration, endovascular stent grafts are considered class III devices[23]. Despite being composed of similar materials to endovascular stent grafts already on the market, the change in geometry will require the fish gill design to be submitted for a PMA. No current devices on the market use a similar design to prevent type 1a endoleaks. To prepare the device for premarket approval, several standards will be closely adhered to. These standards are all recognized by the FDA, which are used for testing endovascular stent grafts.

Table 1: Regulatory and Testing Standards

  - Intellectual Property

While no patent describes a fish gill-like design there are several patents for the overall shape of the stent. Three of the nearly 60,000 patents filed around the world are described below to gauge an understanding of the patent environment.

    - US7407509B2

Status: Expired as of 2024-01-13

Unique Feature: A traditional endovascular stent with an added tube of material spiraling down the outside of the trunk of the stent. This spiraled tube was added to the design to shunt blood flow from the abdominal aortic aneurysm into branch vessels to prevent the aneurysm from expanding. [24]

    - US9597204B2

Status: Active; Expiration Date: 2033-03-10

Unique Feature: A traditional endovascular forked-shaped stent to facilitate blood flow through the aortic aneurysm. The trunk of the stent is designed to be twisted and constricted about a midpoint in the trunk to form an hourglass shape. The stent is designed in this way to assist with rapid deployment of the stent and also allow for the minimization of mechanical stent failure about this axis. [25]

    - EP1810642B1

Status: Active; Expiration Date: 2027-01-19

Unique Feature: This design implements a traditional stent design with the added suprarenal fixation and barbed anchors. The suprarenal fixation is designed to mitigate the possibility of stent migration from occurring and leaving the aneurysm untreated once again. The barbed anchors are designed to create a non-traumatic holding with the existing aortic walls to keep the stent in place and prevent to occurrence of endoleak Type 1. [26]

  - Exogenous Factors

Prior to insertion of the stent graft into the body, the device must be thoroughly cleaned and sterilized to remove any excess materials, toxic or carcinogenic substances, as well as microorganisms. For cleaning, there are multiple steps that must be taken. For cleaning the polymer graft, an ultrasonic bath in 70% IPA 30% H2O is recommended [27]. Cleaning the nitinol stent will require further steps than just rinsing in IPA and water. To remove organic matter from the metal, a bath in specific chemical cleaner would be used. A medical grade cleaner manufacturer, rbp chemicals, supplies a product called MicroClean™ Ti which is made for removing organic matter. The nitinol metal would then follow with a deoxidizing bath in MicroClean™ BS to improve resistance to corrosion [28]. To finish cleaning, the nitinol metal must be polished. A common way to do this is by electropolishing in 70/30 IPA/Sulfuric Acid [28]. However, there has been a more environmentally friendly form of polishing in development. Plasma electrolytic polishing uses a solution of non-toxic ammonium sulphate salt. In [29], a study showed evidence that this method provided sterilization of the polished surface, low surface roughness, shorter processing time and does not have a negative impact on functionality of the nitinol metal.

Following cleaning, sterilization will be required to remove any microorganisms on the device. There were two main criteria chosen when picking a sterilization method. The first was choosing a method that would not alter the stent graft’s mechanical properties, since the unmet clinical need is solved via the physical design of the stent graft. The other goal was to reduce the exposure of the toxic and carcinogenic substances to the body, whether this is from the device itself or by working in a harmful environment. Ethylene Oxide, and other forms of chemical sterilization have been ruled out, due to recent regulatory restrictions of Ethylene Oxide as a sterilization method [30]. Further research determined the best sterilization method given the requirements was electron beam sterilization, or e-beam for short. E-beam sterilization uses focused electrons to sterilize devices as they are scanned by the machine [31].Ethylene Oxide, and other forms of chemical sterilization have been ruled out, due to recent regulatory restrictions of Ethylene Oxide as a sterilization method [30]. Further research determined the best sterilization method given the requirements was electron beam sterilization, or e-beam for short. E-beam sterilization uses focused electrons to sterilize devices as they are scanned by the machine [31].[30]. Further research determined the best sterilization method given the requirements was electron beam sterilization, or e-beam for short. E-beam sterilization uses focused electrons to sterilize devices as they are scanned by the machine [30], [31]. Further research determined the best sterilization method given the requirements was electron beam sterilization, or e-beam for short. E-beam sterilization uses focused electrons to sterilize devices as they are scanned by the machine [31].Ethylene Oxide, and other forms of chemical sterilization have been ruled out, due to recent regulatory restrictions of Ethylene Oxide as a sterilization method [30]. Further research determined the best sterilization method given the requirements was electron beam sterilization, or e-beam for short. E-beam sterilization uses focused electrons to sterilize devices as they are scanned by the machine [31].[30]. Further research determined the best sterilization method given the requirements was electron beam sterilization, or e-beam for short. E-beam sterilization uses focused electrons to sterilize devices as they are scanned by the machine [31]. The pros of using e-beam are reduced cost, faster sterilization than gamma irradiation, negligible emissions, well controlled dose range, minimal atmospheric effects, high dosage rate and sterility assurance level. Most importantly, it does not degrade polymers and protects material properties. This method is an FDA approved and internationally accepted method of sterilization. Some drawbacks to this sterilization method are it is expensive to construct, less penetrative than gamma irradiation and not widely used for bulk sterilization.

  - Achilles Heel

    - Natural Variance

The differences between the same piece of tissue from one person to another varies greatly which can make it difficult to put definitive standards on mechanical properties. The KELOS Endovascular Stent relies on increased pressure within the aneurysm sac to open the unidirectional release valves and allow blood flow back into the lumen of the stent. However, there is no set pressure at which an aneurysm will burst. One aneurysm may grow by a cm and be okay where the next may grow only half that and burst. To ensure that the graft material has the correct pliability and recoil extensive testing would have to be done, and a set open pressure will have to be determined. If the current materials approved by the FDA do not fall within our mechanical property standards, an alternative material will have to be explored. While any single device cannot be a blanket fix due to natural variation, we would hope to set the open pressure low enough, so it opens for 95% of patients that experience endoleak that results in an increased aneurysm sac pressure.

    - Leak Notification

Theoretically, the KELOS stent is supposed to completely relieve the excess pressure created by the leaking graft. If this is the case, the patient should not feel any of the normal symptoms associated with an endovascular leak such as abdominal pain. A lack of symptoms will delay both the doctor and the patient from catching the leak and fixing the defective stent sooner. However, this is not a huge concern for our stent, and is just something to keep in mind. Patients receiving any intervention for an aneurysm should receive regular follow-ups to check for damage to the stent. Since not all types of leak produce symptoms, the KELOS Endovascular Stent should properly reduce pressure within the aneurysm sac until the leak can be found upon regular follow-up.

- References

[1]	“Abdominal Aortic Aneurysm Causes, Symptoms, Treatments.” Accessed: Apr. 08, 2024. [Online]. Available: https://www.uchicagomedicine.org/conditions-services/heart-vascular/aortic-disease/abdominal-aortic-aneurysm

[2]	“Abdominal aortic aneurysm: A comprehensive review - PMC.” Accessed: Apr. 08, 2024. [Online]. Available: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC3076160/

[3]	“Endovascular Stent Graft Repair | Surgery | University of Nebraska Medical Center.” Accessed: Apr. 07, 2024. [Online]. Available: https://www.unmc.edu/surgery/divisions/vascular/patient-care/procedures/endovascular-stent.html

[4]	T. W. Kassem, “Follow up CT angiography post EVAR: Endoleaks detection, classification and management planning,” Egypt. J. Radiol. Nucl. Med., vol. 48, no. 3, pp. 621–626, Sep. 2017, doi: 10.1016/j.ejrnm.2017.03.025.

[5]	T. F. O’Donnell et al., “Late Type 1A Endoleaks: Associated Factors, Prognosis and Management Strategies,” Ann. Vasc. Surg., vol. 80, pp. 273–282, Mar. 2022, doi: 10.1016/j.avsg.2021.08.057.

[6]	T. Duerig, A. Pelton, and D. Stöckel, “An overview of nitinol medical applications,” Mater. Sci. Eng. A, vol. 273–275, pp. 149–160, Dec. 1999, doi: 10.1016/S0921-5093(99)00294-4.

[7]	R. Cassano et al., “Expanded Polytetrafluoroethylene Membranes for Vascular Stent Coating: Manufacturing, Biomedical and Surgical Applications, Innovations and Case Reports,” Membranes, vol. 13, no. 2, p. 240, Feb. 2023, doi: 10.3390/membranes13020240.

[8]	“Coronary & Cardiovascular Stent Tubing - Nitinol & Titanium | Tecomet Etch.” Accessed: Apr. 07, 2024. [Online]. Available: http://www.tecometetch.com/stents.html

[9]	J. Fu, Y. Su, Y.-X. Qin, Y. Zheng, Y. Wang, and D. Zhu, “Evolution of metallic cardiovascular stent materials: a comparative study among stainless steel, magnesium and zinc,” Biomaterials, vol. 230, p. 119641, Feb. 2020, doi: 10.1016/j.biomaterials.2019.119641.

[10]	K. Maleckis et al., “Nitinol Stents in the Femoropopliteal Artery: A Mechanical Perspective on Material, Design, and Performance,” Ann. Biomed. Eng., vol. 46, no. 5, pp. 684–704, May 2018, doi: 10.1007/s10439-018-1990-1.

[11]	“S.M.A.R.T. CONTROLTM Self-Expanding Nitinol Stent,” Cordis. Accessed: Apr. 07, 2024. [Online]. Available: https://cordis.com/apac/products/intervene/endovascular/self-expanding-stents/s-m-a-r-t-control-self-expanding-nitinol-stent

[12]	https://fyra.io, “Stent Graft Material Factors That Impact Limb Complication Rates,” Endovascular Today. Accessed: Apr. 08, 2024. [Online]. Available: https://evtoday.com/articles/2015-sept-supplement/stent-graft-material-factors-that-impact-limb-complication-rates

[13]	“Fig. 1. A. Knitted crimped polyethylene-terephthalate (PET) vascular...,” ResearchGate. Accessed: Apr. 08, 2024. [Online]. Available: https://www.researchgate.net/figure/A-Knitted-crimped-polyethylene-terephthalate-PET-vascular-prostheses-6-mm-with_fig1_259626204

[14]	K. Modi, M. P. Soos, and K. Mahajan, “Stent Thrombosis,” in StatPearls, Treasure Island (FL): StatPearls Publishing, 2024. Accessed: Apr. 08, 2024. [Online]. Available: http://www.ncbi.nlm.nih.gov/books/NBK441908/

[15]	A. Edris, N. Siddiqi, and M. J. Kern, “Restenosis and Drug-Eluting Stents,” in Interventional Cardiac Catheterization Handbook (Third Edition), M. J. Kern, Ed., Philadelphia: W.B. Saunders, 2013, pp. 168–183. doi: 10.1016/B978-0-323-08057-6.00007-0.

[16]	“Neointimal Hyperplasia - StatPearls - NCBI Bookshelf.” Accessed: Apr. 08, 2024. [Online]. Available: https://www.ncbi.nlm.nih.gov/books/NBK499893/

[17]	U. Asenbaum et al., “Stent-graft surface movement after endovascular aneurysm repair: baseline parameters for prediction, and association with migration and stent-graft-related endoleaks,” Eur. Radiol., vol. 29, no. 12, pp. 6385–6395, 2019, doi: 10.1007/s00330-019-06282-w.

[18]	G. A. J. Fransen, P. Desgranges, R. J. F. Laheij, P. L. Harris, J.-P. Becquemin, and EUROSTAR Collaborators, “Frequency, predictive factors, and consequences of stent-graft kink following endovascular AAA repair,” J. Endovasc. Ther. Off. J. Int. Soc. Endovasc. Spec., vol. 10, no. 5, pp. 913–918, Oct. 2003, doi: 10.1177/152660280301000511.

[19]	“Fish - Gills, Lungs, Swim Bladder | Britannica.” Accessed: Apr. 08, 2024. [Online]. Available: https://www.britannica.com/animal/fish/The-respiratory-system

[20]	J. M. Wilson and P. Laurent, “Fish gill morphology: inside out,” J. Exp. Zool., vol. 293, no. 3, pp. 192–213, Aug. 2002, doi: 10.1002/jez.10124.

[21]	“GORE Excluder AAA Endoprosthesis,” Gore Medical. Accessed: Apr. 08, 2024. [Online]. Available: https://www.goremedical.com/products/excluder

[22]	“Alto Abdominal Stent Graft System,” Endologix. [Online]. Available: https://endologix.com/aaa/alto/

[23]	“Product Classification,” U.S. Food & Drug Administration. Accessed: Apr. 02, 2024. [Online]. Available: https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfPCD/classification.cfm?ID=MIH

[24]	R. Greenberg, K. West, D. Skender, and J. Foster, “Branched vessel endoluminal device with fenestration,” US20060247761A1, Nov. 02, 2006 Accessed: Apr. 08, 2024. [Online]. Available: https://patents.google.com/patent/US20060247761/fr

[25]	R. Benary, A. Shalev, and N. S. Nae, “Branched stent-graft system,” US9597204B2, Mar. 21, 2017 Accessed: Apr. 08, 2024. [Online]. Available: https://patents.google.com/patent/US9597204B2/en?oq=US9597204B2

[26]	D. D. Boucher, M. Dugery, K. L. Johnson, I. J. Khan, D. C. Majercak, and S. L. Rush, “Percutaneous endoprosthesis using suprarenal fixation and barbed anchors,” EP1810642B1, Jul. 01, 2009 Accessed: Apr. 08, 2024. [Online]. Available: https://patents.google.com/patent/EP1810642B1/en?oq=EP1810642B1

[27]	“Medical Textile Processing, manufacturing services, laser cutting, medical device cleaning, medical textile cleaning, medical textile cutting, medical textile sterilizing, textile sterilization.” Accessed: Apr. 04, 2024. [Online]. Available: https://www.surgicalmesh.com/processing.htm

[28]	Administrator, “What to Know About Cleaning, Deoxidizing, Passivating and Polishing Medical Implant Devices,” RBP Chemical Technology. Accessed: Apr. 04, 2024. [Online]. Available: https://rbpchemical.com/2019/07/what-to-know-about-medical-implant-devices/

[29]	K. Navickaitė et al., “Plasma Electrolytic Polishing of Nitinol: Investigation of Functional Properties,” Materials, vol. 14, no. 21, p. 6450, Oct. 2021, doi: 10.3390/ma14216450.

30]	L. Friedman, “E.P.A. Sets Limits on Carcinogenic Gas Used to Sterilize Medical Devices,” The New York Times, Mar. 14, 2024. Accessed: Apr. 05, 2024. [Online]. Available: https://www.nytimes.com/2024/03/14/climate/epa-sets-limits-on-ethylene-oxide.html

[31]	“Ethide Laboratories - Electron Beam vs. Gamma Radiation Sterilization,” Ethide Laboratories. Accessed: Apr. 05, 2024. [Online]. Available: https://ethidelabs.com/electron-beam-vs-gamma-radiation-sterilization/